## Slide 1: Non-destructive nuclear measurements, from R&D to applications
Bertrand PEROT, Nuclear Measurement Laboratory
CEA DES IRESNE / DTN SMTA LMN

## Slide 2: Summary
The Nuclear Measurement Laboratory
Photon imaging
Gamma spectroscopy
Passive & active neutron measurements
Neutron and photon activation analysis

## Slide 3: The Nuclear Measurement Laboratory
25-30 permanent staff
5-10 PhD, internships, temporary researchers
> 50 projects or collaborations on both R&D and applications
3
Radiological characterization
        Gamma-ray spectroscopy
        Passive and active neutron measurements
        Photofission
Physical characterization
        Photon radiography and tomography
        Neutron radiography
Elemental characterization
        Thermal neutron activation analysis
        Fast neutron activation analysis
Nuclear fuel cycle
        Uranium mining, fuel fabrication & reprocessing
        Radioactive waste characterization
        Dismantling and decommissioning
Reactors
        Non-Destructive Exams for Jules Horowitz Reactor
        Nuclear accident studies and monitoring
Other applications
        CBRNE threats detection, nuclear material controls
        Waste recycling for the circular economy

## Slide 4: X-ray Imaging
1

## Slide 5: High energy X-ray imaging
Scintillation plate 800 x 600 mm²
(Gadox Gd2O2S - 400 µm)
Varex K15: 9 MV or 15 MV LINAC
(70 or 130 Gy/min at 1 m)
X-ray source (electron LINAC)
Lead
collimator
45° mirror
Phosphor screen
Camera
Bremsstrahlung
X-ray beam
CINPHONIE casemate in CHICADE facility
N. Estre, D. Eck, J-L. Pettier, E. Payan, et al. High-Energy X-Ray Imaging Applied to Non Destructive Characterization of Large Nuclear Waste Drums, DOI:10.1109/ANIMMA.2013.6727987
5
Rotation/elevation bench
(up to 5 tons)
Lead
shield
Radiography
Tomography
Benoit Geslot #04-117

## Slide 6: R&D on high resolution imaging
Objective: high resolution imaging 200 µm
Analysis of small objects and defects (bubbles, cracks…)
Comparative studies of linear detectors
Blur reduction using CNN (convolutional neural network)
  trained with  50 000 pictures with / without blur
  pictures simulated with MODHERATO (LMN code)
M. Maulin et al., Design of a High-Energy and High-Resolution detector for X-ray computed tomography, EPJ Web of Conferences, ANIMMA2021
M. Maulin PhD thesis, CEA, 2022
New (LLNL)
Old GSO (1800 µm) Resolution  800 µm
New GSO (140 µm) + CNN  200 µm

## Slide 7: R&D on dual-energy imaging (1st simulation steps)
Goal: e Vs. Zeff to improve material identification (electronic density vs. effective atomic number)
Means: imaging at 9 MeV vs. 15 MeV (“low energy” LE vs. “high energy” HE)
A. Sardet et al., High-energy dual-energy computed tomography for the characterization of large and thick objects, iCT2025, 14th Conf. on Indus. Computed Tomography, Antwerp, Belgium
Filtered back-projection
Using elemental and matrix calibration
e Vs. Zeff

## Slide 8: TOMIS mobile high-energy tomography
M9 LINAC
Drum
rotating
platform
up to 5T
Beam aperture
LINAC shield
Linear CdWO4 detector
10 mm thick – 1500 mm wide
M9 LINAC 30 Gy/min @ 1 m

## Slide 9: Gamma-Ray Spectroscopy
2

## Slide 10: Gamma-ray spectroscopy
Radionuclide activities:
Compton background
Net surface extraction
Gamma-ray intensity
Measurement time
Cs-137
Sn(E)
T
)
E
(
I
)
E
(
Sn
)
E
(
A
)
E
(
´
´
=
e
High purity
germanium
detectors
Detection efficiency

## Slide 11: Cleaning and dismantling operations
Fast detection of low-level contamination for instance < 1 Bq.cm-2 of uranium on concrete soils
Large ZnS(Ag) and plastic PVT scintillators  Fast alpha and beta detection (within a few min)
Large NaI(Tl) scintillators  More precise activity confirmation (within  15 min)
185 keV, 235U
63 + 92 keV, 234Th (238U chain)
U contamination on a concrete soil (natural U, K, Th background in blue)
NaI(Tl)
gamma
spectra
U contamination
Background
Francisco Salvador #07-130

## Slide 12: NaI(Tl) spectrum analysis based on energy bands
NaI vs. HPGe | HPGe | NaI(Tl)
Acquisition time | 64 h | 15 min
U surface activity | (8.0 ± 0.6) Bq.cm-2 | (8.5 ± 1.2) Bq.cm-2
U enrichment | (0.8 ± 0.1) % | (0.8 ± 0.2) %
F. Salvador et al., Low and high resolution gamma-ray spectroscopy for the characterization of uranium contamination, NIM A 1065 (2024) 169549
Large natural background with significant fluctuations  estimation by multi-linear regression with U, K, Th bands
Two main U bands 235U with 185 keV  ray 238U with 63+92 keV  rays (234Th)
Uranium activity (Bq/cm²)
U enrichment (235U %wgt)
Patent FR2210553
Francisco
Salvador #07-130

## Slide 13: Gamma logging probes for uranium mining
NGRS Probe
Total couting
CeBr3 Probe
Spectrometric probe
L = 697 mm
⌀ 38 mm
NaI
PM
NaI(Tl) 2.4 cm x 3.9 cm
CeBr3 2.0 cm x 9.6 cm
m1
m2
m3
m4
Crystal
PM
m1
m2
m3
m4
m5
m6
m7
m8

KSP Probe
Total couting
NaI(Tl) 3.0 cm x 7.0 cm
LaBr3(Ce) Probe
Spectrometric probe
LaBr3(Ce) 3.81 cm x 7.62 cm
m1
m2
m3
m4
m5
m6
m7
m8
m9
KOBRA
Thomas Marchais #05-34
Arthur Pellet-Rostaing
#10-139
Patents for spectroscopic probes WO2021/122772 A1 + WO2023/031465 A1

## Slide 14: In Situ Leaching Mines  imbalances in 238Unat chain
Secular equilibrium in U chain is broken
  U / Ra  1 due to differential leaching in roll fronts
  Ra / Rn  1 due to radon escape after drilling
U grade over- or underestimation
in total gamma counting probes
Chart: 
| Uranium | Plomb - Bismuth | Radium
Colonne1 | 0.08 | 0.9 | 0.02
Origin of
gamma rays
609 keV (45%)
1764 keV (15%)
1120 keV (15%)
352 keV (36%)
295 keV (18%)
242 keV (7%)
186 keV (3%)
1001 keV (0.8%)
92 keV (4.2%)
Roll front deposits

## Slide 15: New spectroscopic probes and analysis
C[120-170]keV
MCNP simulations
  Compton counts vs. uranium grade (calibration)
  Uncertainties (disequilibrium, density, mineralogy…)
EXP: ORANO CIME calibration station
  Concrete blocks with U grades up to 10 000 ppm
MCNP
MCNP
EXP
1001 keV
C[960-1046]keV
EXP
B. Pérot, T. Marchais, P.-G. Allinei, H. Toubon, Y. Bensedik, R. Goupillou, A. Berland, Development of Gamma Spectroscopic Tools for Uranium Ore Samples and Borehole Exploration, 2022 IEEE Nuclear Science Symposium, 5 – 12 November 2022, Milano, Italy
Patent WO2023/031465 A1
Thomas Marchais #05-34

## Slide 16: New approaches in logging analysis
Bayesian approach (below)
Machine Learning (under progress)
i
Posterior
distributions
Prior
distributions
U grade
U/Ra
U/Ra
U
Arthur Pellet -Rostaing
#10-139

## Slide 17: Neutron Measurements
3

## Slide 18: Main neutron detectors
Passive Non Destructive Assay of Nuclear Materials, Los Alamos National Laboratory, NUREG/CR-5550, 1991
3He gas detectors : golden standard
in radioactive waste characterization, process monitoring, security, SANS and other physics experiments, etc.
But 3He is expensive
Cheap but sensitive to gamma rays and no PSD

## Slide 19
Matrix correction: experimental design with MCNP
Experimental design with 104 MCNP simulations
(Taguchi L72 & L32)
Correlation between Pu signals and matrix indicators
(internal monitors & transmission measurements)
Quentin Ducasse, Cyrille Eleon, Bertrand Perot, Nadia Perot, and Pierre-Guy Allinei, Application of Machine Learning to MICADO Passive and Active Neutron Measurement System for the Characterization of Radioactive Waste Drums, IEEE TRANSACTIONS ON NUCLEAR SCIENCE, VOL. 71, NO. 5, MAY 2024
Pu passive + active signals
Internal monitors
AmBe transmission
Cyrille Eleon & Quentin Ducasse
#07-42

## Slide 20: Matrix effect correction: predictive models
#2
#4
#3
#1
Training dataset based on MCNP simulations
Matrices PE, PVC, Fe…
Densities 0.1 to 0.7 g/cm3
Filling heights 50 to 100%
Regression models
Ordinary Least Squares
Random Forest
Multi-layer Perceptron
Results
Passive coincidence counting
Active neutron interrogation
Cyrille Eleon & Quentin Ducasse
#07-42
Achment CHALIL #7-51 (Bayesian)
+
Q. Ducasse, C. Eleon, B. Perot, N. Perot, P.-G. Allinei, Application of Machine Learning to MICADO Passive and Active Neutron Measurement System for the Characterization of Radioactive Waste Drums, IEEE TRANSACTIONS ON NUCLEAR SCIENCE, VOL. 71, NO. 5, MAY 2024

## Slide 21: Plastic scintillators as an alternative to 3He counters
Wood
Iron
-n (and n-n) coincidences
- coincidences
Accidental coincidences
Rossi-alpha histogram (times of next pulses after a 1st detection)
Vincent Bottau #5-70
Time (ns)
16 PVT scintillators 10 cm ×10 cm ×100 cm
With or without 5 cm lead shield
Metallic and organic matrices in 118 L drums
137Cs, 60Co, 252Cf, AmBe sources, Pu samples
“FASTER” electronics (CNRS)
 ns scale (instead of µs with 3He)
 low accidental level (wrt. 3He)

## Slide 22: Coincidence Multiplicity analysis
 | 252Cf (3.6x104 n/s) |  | AmBe (2.3x105 n/s) |  | 60Co (882 kBq) | 
 | R+A | A | R+A | A | R+A | A
M0 | 3379137 | 4479459 | 15577017 | 16671802 | 53638186 | 57771864
M1 | 857974 | 8310 | 1141198 | 102194 | 4947762 | 989025
M2 | 209237 | 1296 | 59474 | 5790 | 251249 | 81991
M3 | 38098 | 257 | 2312 | 253 | 9079 | 3516
M4 | 4806 | 42 | 67 | 14 | 239 | 97
M5 | 401 | 5 | -9 | 0 | -11 | 2
M6 | 18 | 1 | -3 | 0 | -7 | 0
M7 | -1 | 0 | -2 | 1 | -2 | 0
S | 14 965.6 s-1 |  | 55 933.5 s-1 |  | 196 155.0 s-1 | 
D | 4 667.6 s-1 |  | 3 842.5 s-1 |  | 14 381.3 s-1 | 
T | 1 174.8 s-1 |  | 174.9 s-1 |  | 342.2 s-1 | 
Shift register method to unfold real from accidental coincidences
At least 3 pulses in coincidence to separate spontaneous fissions (Pu or 252Cf) from parasitic coincidences of (,n) reactions, gamma cascades…
Triples  6.7
S, D, T = singles, doubles, triples after real vs. accidental coincidence unfolding
 0.16
V. Bottau, et al, Sorting fission from parasitic coincidences of neutron and gamma rays in plastic scintillators using particle times of flight, ANIMMA 2021 International Conference, Prague, Czech Republic, EPJ Web of Conferences 253, 07014 (2021)
Vincent Bottau #5-70

## Slide 23: Improving S/N ratio using inter-pulse delays
V. Bottau, C. Carasco, B. Perot, C. Eleon, R. De Stefano, L. Isnel, I. Tsekhanovich, Sorting fission from parasitic coincidences of neutron and gamma rays in plastic scintillators using particle times of flight, ANIMMA 2021 International Conference, June 2021, Prague, Czech Republic, EPJ Web of Conferences 253, 07014 (2021)
Radioactive source | Counts in ROI
252Cf (3.6×104 n.s-1) | 10894
AmBe (2.3×105 n.s-1) | 761
60Co (882 kBq) | 522
SNR (252Cf/AmBe) | 14.3*
* After neutron and gamma scattering cross talk rejection (see next slide)
( n n) region of interest
Patents
FR1914402A
(this slide)
FR2406476 (Pu vs. Cm
discrimination)
Vincent Bottau #5-70

## Slide 24: Neutron logging probe for U grade and H porosity
Patent FR2110722
Fission prompt neutrons  U
3He detection block (PE + Cd)
SODERN DT neutron generator
V. Fondement et al., Development of a neutron probe to perform a combined measurement of uranium concentration and hydrogen porosity for mining applications, NIM A 1059 (2024) 168888
Valentin Fondement
#5-143
Backscattered
neutrons  H
H porosity (%)
Count rate (s-1)

## Slide 25: Neutron & photon activation analysis
4

## Slide 26: Pulsed Neutron Activation Analysis
Time
A few 10 µs
Up to few ms
Neutron
flux
Fast neutrons during pulses  (n,n’)
Thermal neutrons between pulses  (n,)
(+ delayed gamma rays of activation)
HPGe gamma spectra
A. Parsons et al., "Subsurface In situ elemental composition measurements with PING," 2013 IEEE Aerospace Conference, Big Sky, MT, USA, 2013
Basalt
Borehole logging (mining, oil..)
Radioactive waste (toxic chemicals, matrix…)
Industry (online analysis, recycling)
Environment (water, soils…)
…

## Slide 27: PGNAA with gamma coincidences
New R&D topics in critical raw material recycling (e.g. detection of dysprosium in permanent magnets)
New approach  (n,) coincidences with large NaI(Tl) scintillators
 Low energy resolution of NaI(Tl) compared to HPGe but high efficiency
 Fast gamma coincidences (ns) to reduce the background and separate gamma rays (see next slide)
PE moderating cell with a
graphite core
5’’x5’’x10’’
NaI(Tl) detectors + Pb and B shield
J. Roult et al. Prompt Gamma Neutron Activation Analysis with Gamma-Gamma Coincidences for Recycling Waste Characterization, 2022 IEEE Nuclear Science Symposium, Milano, Italy
Jeremy Roult dit Rouaux PhD Thesis, Prompt gamma neutron activation analysis for the detection of strategic metals in conventional waste, 2024
One 30 % HPGe
replacing a NaI for comparative tests

## Slide 28: Diagonal projections on 2D spectra
Sulphur sample
Improved signal-to-noise ratio
1D Gaussian fit
gamma pairs whose sum of energies is equal to a specific value : E1 + E2 = 6262 keV
gamma pairs whose sum of energies is equal to a specific value : E1 + E2 = 8642 keV
Jeremy Roult dit Rouaux et al, Journal of Radioanalytical and Nuclear Chemistry (2024) 333:6577–6592 https://doi.org/10.1007/s10967-024-09822-x
Patent WO2025/026740 A1

## Slide 29: The associated particle technique
Unfolding (with a calibration database)
Gamma spectrum (of the selected volume)
Built-in segmented alpha detector
Tagged neutrons
Volume of interest
Volume of interest
Alpha detector
(position sensitive)
Neutron TOF (alpha-gamma coincidence)
Fast neutron interrogation with a DT neutron generator embedding an alpha detector

## Slide 30: Tagged Neutron Inspection System, calibration
Al
C
Cl
Cr
Fe
Cu
N
Na
O
Zn
Mn
Al
C
Ca
F
K
Mg
Ni
P
Pb
S
Si
Database of neutron-induced elemental spectra
C. de Sainte Foy et al., "Optimization of the ENTRANCE Rapidly Relocatable Tagged Neutron Inspection System," 2022 IEEE Nuclear Science Symposium and Medical Imaging Conference (NSS/MIC), Italy, 2022

## Slide 31: Field tests in Rotterdam seaport
40 kg of C4 explosive simulant in a cigarettes cargo, 10 min
To be watched in
presentation mode

## Slide 32: Photofission with delayed particle detection
R&D program for U and Pu characterization in large concrete radioactive waste packages
Delayed neutron counting between LINAC pulses (historical method) but high attenuation in concrete
Delayed gamma spectroscopy after irradiation (HPGe) or between pulses (scintillators LaBr, BGO… )
Gamma
detector
LINAC
Waste package
Photofission delayed  rays
U or Pu photofission
High-E
photons
High-E photons (> 6 MeV)
U or Pu
nucleus
Prompt neutrons
Prompt  rays
Prompt  rays
Delayed neutrons
Delayed neutrons
Delayed  rays
Delayed  rays
Photofission
cross
sections
LINAC
Brems- strahlung
photon
spectrum

## Slide 33: HPGe detection tests of U samples in concrete
LINAC head with its neutron shield
Highly shielded HPGe
Irradiation 2h
HPGe acquisition 2h
Transfer to HPGe
and cooling delay
(< 1 min)
Counts
Energy (keV)
With U sample (17 cm depth)
Active background
Manon Delarue, PhD Thesis, 2020 https://theses.fr/2022GRALY060
M. Delarue et al., New measurements of cumulative photofission yields of 239Pu, 235U and 238U with a 17.5 MeV Bremsstrahlung photon beam and progress toward actinide differentiation, NIM A 1040 (2022) 167259
Benoit Geslot #04-117

## Slide 34: U detection with scintillators in pulsed irradiation
LINAC
LaBr
32°
Concrete matrix
Energy threshold | Net counts from 238U | Background counts | Net counts / background
2 | 544017 826 | 69371 | 8
2.5 | 308041599 | 25434 | 12
3 | 171349429 | 6496 | 26
3.5 | 92585310 | 1824 | 51
4 | 48954 226 | 1184 | 41
4.5 | 24960164 | 1004 | 25
5 | 12357118 | 866 | 14
Energy threshold | Net counts from 238U | Background counts | Net counts / background
2 |  | 69371 | 8
2.5 |  | 25434 | 12
3 |  | 6496 | 26
3.5 |  | 1824 | 51
4 |  | 1184 | 41
4.5 |  | 1004 | 25
5 |  | 866 | 14
C. Carasco et al. Photofission delayed gamma-ray measurements in a large cemented radioactive waste drum during LINAC irradiation, NIM A 1053 (2023) 168360
1 s of
irradiation
1 s of
cooling
1 s of
counting
Photofission delayed gamma-ray spectra
in 10 min acquisitions (200 cycles of 3 s)
LaBr3(Ce) scintillator
U sample
positions
Benoit Geslot #04-117

## Slide 35: Conclusion and Prospects
8

## Slide 36
Conclusion
Nondestructive nuclear measurements
      Fast, cost effective, global assessment of large objects, no secondary waste…
      Radiological, physical and elemental characterization
      Wide range of laboratory or in situ applications
Innovation driven by technology, computing power, AI and data sciences …
    Detectors  CdTe/CZT, LaBr3(Ce), CeBr3… 3He alternatives, gamma and neutron cameras…
    Electronics  digital pulse processing, pulse dating, on board data processing, compact, fast, high dynamic range and resolution…
    Radiation sources  high intensity, low noise, compact, transportable, self-shielded, high energy and spatial resolution (small spot LINAC for imaging), new laser sources (X-ray + neutron) …
    High performance computing  realistic numerical simulations (Monte Carlo neutron and photon transport codes, X-ray imaging…), data processing of both simulation and experimental results (imaging, pulse-by-pulse analysis of list mode data…)
    Data sciences and AI  Sensitivity Analysis, Machine Learning, Bayesian methods…

## Slide 37: Prospects
  Nuclear industry
      Nuclear fuel cycle: uranium mining and enrichment, fuel fabrication, spent fuel reprocessing and recycling
      Reactors: fission (in/ex core), fusion (diagnostics), severe accident monitoring…
      Radioactive waste characterization + Cleaning and Dismantling nuclear facilities
  Non nuclear applications
      Non-destructive testing: online process monitoring, infrastructure inspection, petrol and mining exploration/exploitation, recycling / circular economy …
      Non industrial applications: environment, biology, medicine, physics, archeology …
  Defense, Homeland Security, Safeguards
      Illicit trafficking and CBRNE threats detection: explosives, narcotics, chemical warfare, nuclear materials, contraband goods …
      Safeguards: transportable systems to characterize nuclear materials

## Slide 38
bertrand.perot@cea.fr
Thank you!
CEA Cadarache